| DATE OF ISSUE: | 17.09.2026 |
| --- | --- |
| DATE OF CALIBRATION: | 10.09.- 17.09.2026 |
| CUSTOMER: | {{ pjr_customer_name }} |
|  | {{ pjr_customer_address_city }} |
|  | {{ pjr_customer_address_street }} |
|  | {{ pjr_customer_address_state }} |
| CALIBRATION FACILITY: | INTRIPLE, a.s. |
|  | K Zelenči 2957/5 |
|  | 19300 Praha |
|  | Czech Republic |
| EQUIPMENT: | IZ245 – Monitoring receiver |
| OPTION: |  |
| SERIAL NUMBER: | {{ pjr_serial_number }} |
| The expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k=2, which for a normal distribution corresponds to a coverage probability of approximately 95%, unless otherwise stated. This calibration certificate documents, that the named item is tested and measured against defined specifications. Calibration is performed on test equipment and according to standards that are directly or indirectly traceable using approved calibration techniques according to CMI or other national/international standards. The standards implement physical units of measurement according to the International System of Units (SI). If no standards are available, measurements are always based on the relevant standards of INTRIPLE, a.s. The principles and methods of calibration comply with and are in accordance with EN ISO/IEC 17025. The quality system used is certified according to EN ISO 9001. Calibration certificates may not be reproduced other than in full. Calibration certificates without signatures are not valid. The user is obliged to have the item calibrated at appropriate intervals. The interval is determined by the user's responsible person. |  |

| RESULTS: | All measurement results are within specification |  |
| --- | --- | --- |
| MEASURED BY: | {{ pjr_measured_by }} |  |
| TEMPERATURE: | 24°C ±4°C |  |
| HUMIDITY: | 20% to 70% |  |
|  |  | Ing. Michal Groh |
|  |  | Head of Laboratory |

Contents:

2	TEST EQUIPMENT SUMARY	3

3	HARDWARE AND FIRMWARE CONFIGURATION	4

4	Test Summary	4

5	CALIBRATION	5

5.1	TEST1 – CALIBRATION VERIFICATION – RF1 INPUT	5

5.2	TEST1– CALIBRATION VERIFICATION – RF2 INPUT	16

5.3	TEST 2 – NOISE LEVEL MEASUREMENT	18

5.4	TEST 3 – PRESELECTOR	22

5.5	TEST 4 – INTERNAL TEST SYNTHESIZER	27

5.6	TEST 5 – PHASE NOISE	28

5.7	TEST 6 – AMPLITUDE NOISE	31

5.8	TEST 7 – IP3 AND SFDR MEASUREMENT	34

5.9	TEST 9 – IF FILTER FLATNESS	35

5.10	TEST 10 – REFERENCE OSCILLATOR ACCURACY	37

# TEST EQUIPMENT SUMARY

Supplier fa. INTRIPLE a.s. performed final inspection of individual components and entire system to meet the guaranteed technical parameters before sending the measurement system. For the measurements are used company measuring devices, which are regularly calibrated. List of used measuring instruments:

| DEVICE: | MANUFACTURER: | DETAIL: |
| --- | --- | --- |
| SMA100B – Signal generator | Rohde & Schwarz | Type; Signal generator. Frequency Range; 8 kHz – 20 GHz. Serial Number; 103462. Calibration Due; 5.3.2027. |
| SMA100B – Signal generator | Rohde & Schwarz | Type; Signal generator. Frequency Range; 8 kHz – 20 GHz. Serial Number; 101406. Calibration Due; 14.11.2028. |
| 53230A – Frequency counter | Agilent | Type; Frequency counter. Frequency Range; 10 Hz - 350 MHz. Serial Number; MY50001845. Calibration Due; 4.5.2029. |
| NRP18TN | Rohde & Schwarz | Type; Power Meter. Frequency Range; DC - 18 GHz. Serial Number; 101293. Calibration Due; 5.8.2028. |
| MSO7104B | Agilent | Type; Oscilloscope. Frequency Range; 1 GHz. Serial Number; MY50340815. Calibration Due; 5.3.2027. |
| CBL-4FT-SMSM+ | Mini-Circuits | Type; Cable. Serial Number; 223111. |
| BW-S3W2+ | Mini-Circuits | Type; Attenuator. Serial Number; 2027. |
| BW-S3W2+ | Mini-Circuits | Type; Attenuator. Serial Number; 2045. |
| BW-S3W2+ | Mini-Circuits | Type; Attenuator. Serial Number; 2027. |

# HARDWARE AND FIRMWARE CONFIGURATION

| HARDWARE VERSION RF BOARD: | {{ ver_hw_rf}} |
| --- | --- |
| HARDWARE VERSION FPGA BOARD: | {{ ver_hw_fpga}} |
| FPGA VERSION: | {{ ver_fpga}} |
| FIRMWARE VERSION: | {{ ver_fw}} |

# Test Summary

| Name | Description of test | Result | Test date |
| --- | --- | --- | --- |
| Test 1 | Calibration Verification | PASS | {{ test1_date}} |
| Test 2 | Noise Level | PASS | {{ test2_date}} |
| Test 3 | Preselector | PASS | {{ test3_date}} |
| Test 4 | Test Synthesizer | PASS | {{ test4_date}} |
| Test 5 | Phase Noise | PASS | {{ test5_date}} |
| Test 6 | Amplitude Noise | PASS | {{ test6_date}} |
| Test 7 | IP3 Measurement | PASS | {{ test7_date}} |
| Test 9 | IF filter shape | PASS | {{ test8_date}} |
| Test 10 | Reference oscillator | PASS | {{ test9_date}} |

# CALIBRATION

| SERIAL NUMBER: | {{ serial_number}} |
| --- | --- |
| FREQUENCY RANGE: | 1 MHz – 8 GHz |
| TESTED ITEM: | IZ245 |
| CALIBRATION PROCEDURE: | KP18.1 |
| RECEIVER TEMPERATURE: | {{ calibration_temperature}} |
| RESULT: | PASS |

An expanded measurement uncertainty is ± 2 dB. The expanded uncertainty with a coverage factor k = 2 corresponds to a confidence level of approximately 95%.